## Calibration
MagVenture C-B60 ¦ MagPro X100
biphasic ¦ current direction: normal ¦ runtime curve: 0-100% MSO
amplitude_pct_mso | realized_didt_A_per_us | secondary_phase_didt_A_per_us | coil_temp_C | measurement_status | source_image | note | linear_fit_A_per_us | Measured-only slope | 1.6043
0 | 0 |  |  | physical_zero_anchor |  | Physical zero-output anchor | -2.9748 | Measured-only intercept | -2.9748
1 | 1.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | -1.3705 | Measured-only R-squared | 0.99973
2 | 2.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 0.23376 | Minimum runtime MSO | 0
3 | 4.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 1.838 | Maximum runtime MSO | 100
4 | 5.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 3.4423 | Minimum source-observation MSO | 5
5 | 7 |  | 23 | inferred_source_amplitude | IMG_1546 2.jpeg | Display photographed at 0% after the pulse; 7 A/us assigned to the intended preceding 5% protocol step | 5.0466 | Maximum source-observation MSO | 80
6 | 8.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 6.6509
7 | 9.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 8.2551
8 | 10.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 9.8594
9 | 11.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 11.4637
10 | 13 |  | 23 | measured | IMG_1547 2.jpeg |  | 13.068
11 | 14.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 14.6722
12 | 16.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 16.2765
13 | 17.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 17.8808
14 | 19.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 19.4851
15 | 21 |  | 23 | measured | IMG_1548 2.jpeg |  | 21.0893
16 | 22.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 22.6936
17 | 24.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 24.2979
18 | 25.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 25.9022
19 | 27.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 27.5064
20 | 29 |  | 23 | interpolated |  | Linear interpolation between measured 15% and 25% rows | 29.1107
21 | 30.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 30.715
22 | 32.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 32.3193
23 | 33.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 33.9235
24 | 35.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 35.5278
25 | 37 |  | 23 | measured | IMG_1549 2.jpeg |  | 37.1321
26 | 38.7 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 38.7364
27 | 40.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 40.3407
28 | 42.1 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 41.9449
29 | 43.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 43.5492
30 | 45.5 |  | 23 | interpolated |  | Linear interpolation between measured 25% and 35% rows | 45.1535
31 | 47.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 46.7578
32 | 48.9 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 48.362
33 | 50.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 49.9663
34 | 52.3 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 51.5706
35 | 54 |  | 23 | measured | IMG_1550 2.jpeg |  | 53.1749
36 | 55.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 54.7791
37 | 56.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 56.3834
38 | 58.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 57.9877
39 | 59.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 59.592
40 | 61 |  | 23 | measured | IMG_1551 2.jpeg |  | 61.1962
41 | 62.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 62.8005
42 | 64.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 64.4048
43 | 66.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 66.0091
44 | 68.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 67.6133
45 | 70 |  | 23 | measured | IMG_1552 2.jpeg |  | 69.2176
46 | 71.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 70.8219
47 | 72.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 72.4262
48 | 74.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 74.0304
49 | 75.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 75.6347
50 | 77 |  | 23 | measured | IMG_1553 2.jpeg; IMG_1554 2.jpeg | Duplicate photographs agree | 77.239
51 | 78.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 78.8433
52 | 79.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 80.4475
53 | 81.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 82.0518
54 | 82.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 83.6561
55 | 84 |  | 23 | measured | IMG_1555 2.jpeg |  | 85.2604
56 | 85.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 86.8646
57 | 87.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 88.4689
58 | 89.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 90.0732
59 | 91.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 91.6775
60 | 93 |  | 23 | measured | IMG_1556 2.jpeg |  | 93.2817
61 | 94.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 94.886
62 | 96.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 96.4903
63 | 97.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 98.0946
64 | 99.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 99.6989
65 | 101 |  | 23 | measured | IMG_1557 2.jpeg |  | 101.3031
66 | 102.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 102.9074
67 | 104.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 104.5117
68 | 106.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 106.116
69 | 108.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 107.7202
70 | 110 |  | 23 | measured | IMG_1558 2.jpeg |  | 109.3245
71 | 111.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 110.9288
72 | 113.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 112.5331
73 | 114.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 114.1373
74 | 116.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 115.7416
75 | 118 |  | 23 | measured | IMG_1559 2.jpeg |  | 117.3459
76 | 119.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 118.9502
77 | 120.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 120.5544
78 | 122.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 122.1587
79 | 123.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 123.763
80 | 125 |  | 23 | measured | IMG_1561 2.jpeg |  | 125.3673
81 | 126.6043 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 126.9715
82 | 128.2086 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 128.5758
83 | 129.8128 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 130.1801
84 | 131.4171 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 131.7844
85 | 133.0214 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 133.3886
86 | 134.6257 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 134.9929
87 | 136.2299 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 136.5972
88 | 137.8342 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 138.2015
89 | 139.4385 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 139.8057
90 | 141.0428 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 141.41
91 | 142.647 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 143.0143
92 | 144.2513 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 144.6186
93 | 145.8556 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 146.2228
94 | 147.4599 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 147.8271
95 | 149.0641 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 149.4314
96 | 150.6684 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 151.0357
97 | 152.2727 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 152.6399
98 | 153.877 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 154.2442
99 | 155.4812 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 155.8485
100 | 157.0855 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 80% observation | 157.4528

## Metadata
field | value
id | magventure_x100_c-b60_biphasic_normal
manufacturer | MagVenture
stimulator_model | MagPro X100
coil_hardware_name | MagVenture C-B60
console_coil_type | C-B60
simnibs_coil_model_basename | MagVenture_C-B60.ccd
waveform | biphasic
current_direction | normal
csv | magventure_x100_c-b60_biphasic_normal/didt_calibration.csv
value_column | realized_didt_A_per_us
mso_range | [0.0, 100.0]
measurement_date | 2026-09-10
intended_roles | ["sm_hand_motor_threshold"]
status | measured_and_estimated_full_curve
full_curve_mso_range | [0.0, 100.0]
estimated_above_mso | 80
measured_only_linear_slope_A_per_us_per_pct_mso | 1.6043
measured_only_linear_intercept_A_per_us | -2.9748
source_observation_mso_range | [5.0, 80.0]
canonical_csv | didt_calibration.csv
workbook_role | Derived human-review artifact; CSV is authoritative