## Calibration
MagVenture Cool-B65 A/P ¦ MagPro X100
biphasic ¦ current direction: normal ¦ runtime curve: 0-100% MSO
amplitude_pct_mso | realized_didt_A_per_us | secondary_phase_didt_A_per_us | coil_temp_C | measurement_status | source_image | note | linear_fit_A_per_us | Measured-only slope | 1.549
0 | 0 |  |  | physical_zero_anchor |  | Physical zero-output anchor | -3.1176 | Measured-only intercept | -3.1176
1 | 0.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | -1.5686 | Measured-only R-squared | 0.99985
2 | 1.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | -0.019608 | Minimum runtime MSO | 0
3 | 2.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 1.5294 | Maximum runtime MSO | 100
4 | 3.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 3.0784 | Minimum source-observation MSO | 5
5 | 4 |  | 24 | measured | IMG_1569 2.jpeg |  | 4.6275 | Maximum source-observation MSO | 85
6 | 5.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 6.1765
7 | 7.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 7.7255
8 | 9.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 9.2745
9 | 11.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 10.8235
10 | 13 |  | 24 | measured | IMG_1570 2.jpeg |  | 12.3725
11 | 14.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 13.9216
12 | 15.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 15.4706
13 | 17.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 17.0196
14 | 18.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 18.5686
15 | 20 |  | 24 | measured | IMG_1571 2.jpeg |  | 20.1176
16 | 21.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 21.6667
17 | 23.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 23.2157
18 | 24.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 24.7647
19 | 26.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 26.3137
20 | 28 |  | 24 | measured | IMG_1572 2.jpeg |  | 27.8627
21 | 29.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 29.4118
22 | 30.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 30.9608
23 | 32.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 32.5098
24 | 33.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 34.0588
25 | 35 |  | 24 | measured | IMG_1573 2.jpeg |  | 35.6078
26 | 36.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 37.1569
27 | 38.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 38.7059
28 | 40.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 40.2549
29 | 42.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 41.8039
30 | 44 |  | 24 | measured | IMG_1574 2.jpeg |  | 43.3529
31 | 45.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 44.902
32 | 46.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 46.451
33 | 48.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 48
34 | 49.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 49.549
35 | 51 |  | 24 | measured | IMG_1575 2.jpeg |  | 51.098
36 | 52.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 52.6471
37 | 54.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 54.1961
38 | 55.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 55.7451
39 | 57.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 57.2941
40 | 59 |  | 24 | measured | IMG_1576 2.jpeg |  | 58.8431
41 | 60.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 60.3922
42 | 61.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 61.9412
43 | 63.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 63.4902
44 | 64.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 65.0392
45 | 66 |  | 24 | measured | IMG_1577 2.jpeg |  | 66.5882
46 | 67.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 68.1373
47 | 69.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 69.6863
48 | 71.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 71.2353
49 | 73.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 72.7843
50 | 75 |  | 24 | measured | IMG_1578 2.jpeg |  | 74.3333
51 | 76.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 75.8824
52 | 77.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 77.4314
53 | 79.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 78.9804
54 | 80.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 80.5294
55 | 82 |  | 24 | measured | IMG_1579 2.jpeg |  | 82.0784
56 | 83.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 83.6275
57 | 85.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 85.1765
58 | 86.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 86.7255
59 | 88.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 88.2745
60 | 90 |  | 24 | measured | IMG_1580 2.jpeg |  | 89.8235
61 | 91.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 91.3725
62 | 92.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 92.9216
63 | 94.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 94.4706
64 | 95.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 96.0196
65 | 97 |  | 24 | measured | IMG_1581 2.jpeg |  | 97.5686
66 | 98.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 99.1176
67 | 100.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 100.6667
68 | 102.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 102.2157
69 | 104.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 103.7647
70 | 106 |  | 24 | measured | IMG_1582 2.jpeg |  | 105.3137
71 | 107.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 106.8627
72 | 108.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 108.4118
73 | 110.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 109.9608
74 | 111.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 111.5098
75 | 113 |  | 24 | measured | IMG_1583 2.jpeg |  | 113.0588
76 | 114.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 114.6078
77 | 116.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 116.1569
78 | 117.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 117.7059
79 | 119.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 119.2549
80 | 121 |  | 24 | measured | IMG_1584 2.jpeg |  | 120.8039
81 | 122.4 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 122.3529
82 | 123.8 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 123.902
83 | 125.2 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 125.451
84 | 126.6 |  |  | interpolated |  | Piecewise-linear interpolation between observed points | 127
85 | 128 |  | 24 | measured | IMG_1585 2.jpeg |  | 128.549
86 | 129.549 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 130.098
87 | 131.098 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 131.6471
88 | 132.6471 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 133.1961
89 | 134.1961 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 134.7451
90 | 135.7451 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 136.2941
91 | 137.2941 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 137.8431
92 | 138.8431 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 139.3922
93 | 140.3922 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 140.9412
94 | 141.9412 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 142.4902
95 | 143.4902 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 144.0392
96 | 145.0392 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 145.5882
97 | 146.5882 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 147.1373
98 | 148.1373 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 148.6863
99 | 149.6863 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 150.2353
100 | 151.2353 |  |  | model_extrapolated |  | Measured-only linear slope extrapolation anchored to the 85% observation | 151.7843

## Metadata
field | value
id | magventure_x100_cool-b65-ap_biphasic_normal
manufacturer | MagVenture
stimulator_model | MagPro X100
coil_hardware_name | MagVenture Cool-B65 A/P
console_coil_type | Cool-B65-A/P
simnibs_coil_model_basename | MagVenture_Cool-B65.ccd
waveform | biphasic
current_direction | normal
csv | magventure_x100_cool-b65-ap_biphasic_normal/didt_calibration.csv
value_column | realized_didt_A_per_us
mso_range | [0.0, 100.0]
measurement_date | 2026-09-10
intended_roles | ["treatment"]
study_context | ["BD2"]
status | measured_and_estimated_full_curve
full_curve_mso_range | [0.0, 100.0]
estimated_above_mso | 85
measured_only_linear_slope_A_per_us_per_pct_mso | 1.549
measured_only_linear_intercept_A_per_us | -3.1176
source_observation_mso_range | [5.0, 85.0]
canonical_csv | didt_calibration.csv
workbook_role | Derived human-review artifact; CSV is authoritative