## Calibration
MagVenture Cool-B65 ¦ MagPro X100
biphasic ¦ current direction: normal ¦ runtime curve: 0-100% MSO
amplitude_pct_mso | realized_didt_A_per_us | secondary_phase_didt_A_per_us | coil_temp_C | measurement_status | source_image | note | linear_fit_A_per_us | Measured-only slope | 1.5348
0 | 0 | 0 |  | physical_zero_anchor |  | Physical zero-output anchor | -3.0379 | Measured-only intercept | -3.0379
1 | 0.8 | 1.3 |  | interpolated |  | Piecewise-linear interpolation between observed points | -1.5031 | Measured-only R-squared | 0.99981
2 | 1.6 | 2.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 0.031621 | Minimum runtime MSO | 0
3 | 2.4 | 3.9 |  | interpolated |  | Piecewise-linear interpolation between observed points | 1.5664 | Maximum runtime MSO | 100
4 | 3.2 | 5.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 3.1011 | Minimum source-observation MSO | 5
5 | 4 |  | 23 | measured | IMG_1564 2.jpeg | 2026-09-10 revalidation point | 4.6359 | Maximum source-observation MSO | 100
6 | 5.8 | 7.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 6.1707
7 | 7.6 | 9.1 |  | interpolated |  | Piecewise-linear interpolation between observed points | 7.7054
8 | 9.4 | 10.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 9.2402
9 | 11.2 | 11.7 |  | interpolated |  | Piecewise-linear interpolation between observed points | 10.775
10 | 13 | 13 | 23 | measured | IMG_1565 2.jpeg | 2026-09-10 revalidation point; replaces legacy phase-A value 14 A/us | 12.3097
11 | 14.6 | 14.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 13.8445
12 | 16.2 | 16.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 15.3793
13 | 17.8 | 17.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 16.914
14 | 19.4 | 19.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 18.4488
15 | 21 | 21 | 25 | measured | legacy_calibration_session |  | 19.9835
16 | 22.4 | 22.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 21.5183
17 | 23.8 | 23.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 23.0531
18 | 25.2 | 25.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 24.5878
19 | 26.6 | 26.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 26.1226
20 | 28 | 28 | 25 | measured | legacy_calibration_session |  | 27.6574
21 | 29.4 | 29.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 29.1921
22 | 30.8 | 30.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 30.7269
23 | 32.2 | 32.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 32.2616
24 | 33.6 | 33.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 33.7964
25 | 35 | 35 | 25 | measured | legacy_calibration_session |  | 35.3312
26 | 36.4 | 36.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 36.8659
27 | 37.8 | 38.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 38.4007
28 | 39.2 | 40.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 39.9355
29 | 40.6 | 42.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 41.4702
30 | 42 | 44 | 23 | measured | IMG_1566 2.jpeg | 2026-09-10 revalidation point; replaces legacy phase-A value 44 A/us | 43.005
31 | 43.8 | 45.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 44.5398
32 | 45.6 | 46.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 46.0745
33 | 47.4 | 48.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 47.6093
34 | 49.2 | 49.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 49.144
35 | 51 | 51 | 25 | measured | legacy_calibration_session |  | 50.6788
36 | 52.4 | 52.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 52.2136
37 | 53.8 | 54.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 53.7483
38 | 55.2 | 55.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 55.2831
39 | 56.6 | 57.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 56.8179
40 | 58 | 59 | 25 | measured | legacy_calibration_session |  | 58.3526
41 | 59.6 | 60.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 59.8874
42 | 61.2 | 61.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 61.4221
43 | 62.8 | 63.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 62.9569
44 | 64.4 | 64.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 64.4917
45 | 66 | 66 | 23 | measured | IMG_1567 2.jpeg | 2026-09-10 revalidation point; agrees with legacy phase-A value | 66.0264
46 | 67.4 | 67.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 67.5612
47 | 68.8 | 69.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 69.096
48 | 70.2 | 71.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 70.6307
49 | 71.6 | 73.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 72.1655
50 | 73 | 75 | 26 | measured | legacy_calibration_session |  | 73.7002
51 | 74.8 | 76.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 75.235
52 | 76.6 | 77.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 76.7698
53 | 78.4 | 79.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 78.3045
54 | 80.2 | 80.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 79.8393
55 | 82 | 82 | 26 | measured | legacy_calibration_session |  | 81.3741
56 | 83.4 | 83 |  | interpolated |  | Piecewise-linear interpolation between observed points | 82.9088
57 | 84.8 | 84 |  | interpolated |  | Piecewise-linear interpolation between observed points | 84.4436
58 | 86.2 | 85 |  | interpolated |  | Piecewise-linear interpolation between observed points | 85.9784
59 | 87.6 | 86 |  | interpolated |  | Piecewise-linear interpolation between observed points | 87.5131
60 | 89 | 87 | 23 | measured | IMG_1568 2.jpeg | 2026-09-10 revalidation point; agrees with legacy phase-A value | 89.0479
61 | 90.4 | 88.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 90.5826
62 | 91.8 | 89.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 92.1174
63 | 93.2 | 90.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 93.6522
64 | 94.6 | 91.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 95.1869
65 | 96 | 93 | 26 | measured | legacy_calibration_session |  | 96.7217
66 | 97.6 | 93.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 98.2565
67 | 99.2 | 94.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 99.7912
68 | 100.8 | 95.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 101.326
69 | 102.4 | 96.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 102.8607
70 | 104 | 97 | 26 | measured | legacy_calibration_session |  | 104.3955
71 | 105.8 | 97.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 105.9303
72 | 107.6 | 97.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 107.465
73 | 109.4 | 98.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 108.9998
74 | 111.2 | 98.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 110.5346
75 | 113 | 99 | 26 | measured | legacy_calibration_session |  | 112.0693
76 | 114.5 | 99.7 |  | interpolated |  | Piecewise-linear interpolation between observed points | 113.6041
77 | 116 | 100.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 115.1389
78 | 117.5 | 101.1 |  | interpolated |  | Piecewise-linear interpolation between observed points | 116.6736
79 | 119 | 101.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 118.2084
80 | 120.5 | 102.5 | 27 | interpolated | legacy_calibration_session | Linear interpolation between 75% and 85%; original 80% display image duplicated the 75% state | 119.7431
81 | 122 | 103.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 121.2779
82 | 123.5 | 103.9 |  | interpolated |  | Piecewise-linear interpolation between observed points | 122.8127
83 | 125 | 104.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 124.3474
84 | 126.5 | 105.3 |  | interpolated |  | Piecewise-linear interpolation between observed points | 125.8822
85 | 128 | 106 | 27 | measured | legacy_calibration_session |  | 127.417
86 | 129.4 | 106.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 128.9517
87 | 130.8 | 106.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 130.4865
88 | 132.2 | 107.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 132.0212
89 | 133.6 | 107.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 133.556
90 | 135 | 108 | 27 | measured | legacy_calibration_session |  | 135.0908
91 | 136.4 | 109 |  | interpolated |  | Piecewise-linear interpolation between observed points | 136.6255
92 | 137.8 | 110 |  | interpolated |  | Piecewise-linear interpolation between observed points | 138.1603
93 | 139.2 | 111 |  | interpolated |  | Piecewise-linear interpolation between observed points | 139.6951
94 | 140.6 | 112 |  | interpolated |  | Piecewise-linear interpolation between observed points | 141.2298
95 | 142 | 113 | 27 | measured | legacy_calibration_session |  | 142.7646
96 | 143.8 | 113.8 |  | interpolated |  | Piecewise-linear interpolation between observed points | 144.2994
97 | 145.6 | 114.6 |  | interpolated |  | Piecewise-linear interpolation between observed points | 145.8341
98 | 147.4 | 115.4 |  | interpolated |  | Piecewise-linear interpolation between observed points | 147.3689
99 | 149.2 | 116.2 |  | interpolated |  | Piecewise-linear interpolation between observed points | 148.9036
100 | 151 | 117 | 27 | measured | legacy_calibration_session |  | 150.4384

## Metadata
field | value
id | magventure_x100_cool-b65_biphasic_normal
manufacturer | MagVenture
stimulator_model | MagPro X100
coil_hardware_name | MagVenture Cool-B65
console_coil_type | Cool-B65
simnibs_coil_model_basename | MagVenture_Cool-B65.ccd
waveform | biphasic
current_direction | normal
csv | magventure_x100_cool-b65_biphasic_normal/didt_calibration.csv
value_column | realized_didt_A_per_us
secondary_phase_column | secondary_phase_didt_A_per_us
mso_range | [0.0, 100.0]
measurement_date
revalidation_date | 2026-09-10
intended_roles | ["treatment"]
study_context | ["BioSelect"]
status | measured_and_estimated_full_curve
full_curve_mso_range | [0.0, 100.0]
estimated_above_mso
measured_only_linear_slope_A_per_us_per_pct_mso | 1.5348
measured_only_linear_intercept_A_per_us | -3.0379
source_observation_mso_range | [5.0, 100.0]
canonical_csv | didt_calibration.csv
workbook_role | Derived human-review artifact; CSV is authoritative